| BOARD NAME: | CCD 8 CH Acquisition Board |  | Next Asbly: |  |  | MNSN-EL-10-0500-029 |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| BRD SERL # | 029 |  |  |  | MNSN-EL-07-0501-029 |  |  |  |  |
| PN# | MNSN-EL-10-0500-029 |  |  |  |  |  |  |  |  |
| ASBLY# | MNSN-EL-04-0002 |  |  | PCB# | MNSN-EL-04-1002 |  |  | REV | B |
| BOM# | MNSN-EL-04-4002 |  |  | SCH# | MNSN-EL-04-2002 |  |  | REV | B |
| COGNIZANT ENGNR |  |  |  |  |  |  |  |  |  |

| DATE | SUB-ASSEMBLY OPERATION RECORD | SRL # / ECO # CHECKSUM | NAME |
| --- | --- | --- | --- |
| 03/02/06 | Create ART’s, take boards down for inspection Relevant documents: MNSN-EL-04-0002_rB, MNSN-EL-04-2002_rB, MNSN-EL-04-4002_rB2 |  | dms |
| 3/29/06 | U127 pin 6 not soldered,and pin 1 also. U178 next to it in C397 not soldered. Did not check voltage Did not put serial #on bound |  |  |
| 05/26/2006 | Removed F2, Jumper between F2 and F6 (TP32 – TP36), Installed Resistors 0 Ohm 1206, R168 = 0, R170 = 0, R180 = 0, R183 = 0, R186 = 0, R190 = 0, R238, R239 = 0, R274 = 0, R275 = 0, R633 = 0, R634 = 0,R671 = 0,R678 Replaced Resistors 0805, R264 = 649, R270 = 649, R271 = 1.65K R730 = 2.32K, R740 = 2.32K. | MNSN-0118 | V.A.C |
| 05/26/2006 | Missing C397, but we installed momentarily a capacitor of a similar characteristic, because we don’t have the replacement. Pin 1 and 6 U127 ,repaired Pin 3 U119 not soldered , repaired |  | V.A.C |
| 06/01/2006 | Channel 1 and Channel 3 have replacement and inferior integration caps fitted (C397 & C458).These must be replaced before shipment, original board had caps miss mounted. Functional Test ,File Name is. MNSN-EL-08-0500-SN029-001.btr |  | V.A.C |
| 06/08/2006 | Update the firmware, with a new version v411 checksum and usercode. Modified integrate circuit , pin 2 of U7,U18,U30,U42,U77,U88,U100 and U112. Retested telemetry and video | 0014ffdba 06060202 MNSN-0115 | V.A.C |
| 06/29/06 | Replace C397 & C458 with 330pf capacitor – ECHU1H331JX5 |  | KWD |
| 07/05/06 | Test in system – pass Vendor Board Number 060224-002 |  | KWD |
| 08/31/06 | Assigned to QUOTA |  | dms |
| 09/19/06 | Update Firmware to v491 | 0014c54f7 06090701 | KWD |
| 10/19/06 | Dave Sawyer reports board inop – outputs all zero Board in for recheck |  | KWD |
| 05/22/07 | Problem: Video Ch 1 not functioning, +Vin to ADC at –1V, replaced parts that were removed for troubleshoot: U77,U78,D9,R418,U88,U89, D10, R475; found R416 to be ~2K, should be 10Ω – replaced. Then R418 had unsoldered foot – repaired. OK now; board needs retest. |  | JAC |
| 05/29/07 | Verified Standard Firmware: U143: File – CcdAcqFpgaV411.mcs U68: File – CcdAcqFpgaV411.bit U66 & U67: File CcdAcqCpldV41.jed | 0014ffdba 06060202 a101 | JAC |
| 08/06/09 | Updated board per ECO MNSN-0120 | MNSN-0120 | KZ |
| 08/06/09 | Tested - Passed |  | R.G. |
| 08/06/09 | Installed into 1.5m Echelle system. |  | R.G. |
| 02/11/11 | Modified to be spare for Chiron system: R740, R270, R264 = 2.32K. R731, R263, R271 = 4.00K. R730=653. R742=1.65K. |  | PCM. |