| BOARD NAME: | Clock & Bias Board CCD TYPE |  | Next Asbly: |  |  | MNSN-EL-10-0300 |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| BRD SERL # | 033 |  |  | ART# | MNSN-EL-07-0301-033 |  |  |  |  |
| PN# | MNSN-EL-10-0300-017 |  |  | VND# | 072707-002 |  |  |  |  |
| ASBLY# | MNSN-EL-04-0008 |  |  | PCB# | MNSN-EL-04-1008 |  |  | REV | A |
| BOM# | MNSN-EL-04-4008 |  |  | SCH# | MNSN-EL-04-2008 |  |  |  | A |
| COGNIZANT ENGNR |  | Peter Moore |  |  |  |  |  |  |  |

| DATE | SUB-ASSEMBLY OPERATION RECORD | SRL # / ECO # CHECKSUM | NAME |
| --- | --- | --- | --- |
| 08/02/2007 | Create ART’s, take boards down for inspection. Already has relevant documents: Asbly: MNSN-EL-04-0008_rA, Sch: MNSN-EL-04-2008_rA2.sch PL: MNSN-EL-04-4008-017_rA3.xls |  | dms |
| 08/14/07 | Assign S/N 033 to Board Missing from board U4, U7, U8, U11, U12, U15, U16, U19, U20, U23, U25, U27, U29, U32, U34, U36, U38, U40, U120, U125, U129, U130, U133, U134, U137, U138, U141, U143, U145, U147, U149, U512, U154, U156, U158, U160 INA118UB Missing U100, U169 MAX6143AASA25 INSPECTION COMPLETE INSTALL JUMPERS JP1 – 2,3: JP2 – 1,2: JP3 – 2,3: JP4 – 2,3: JP5 – 2,3: JP6 – 2,3: JP7 – SHORT: JP8 – SHORT: JP9 – 1,2: JP10 – OPEN: JP11 – 1,2 |  | KWD |
| 08/24/07 | Install MAX6143AASA25 in U100, U160 |  | KWD |
| 08/24/07 | R266, R269, R272, R275, R639, R642, R645, R648 – 0 OHM RESISTOR INSTALLED. REPLACE WITH 49.9 OHM RESISTOR |  | JAC |
| 08/27/07 | R260, R261, R262, R235, R613, R512, R614, R615 – 0 OHM RESISTOR INSTALLED REPLACE WITH 49.9 OHM RESISTOR |  | JAC |
| 08/27/07 | LOAD FIRMWARE CLKBIASFPGAv42.mcs CLKBIASFPGAv42.bit CLKBIASCFGCPLDv42.jed CLKBIASSEQCPLDv401.jed | 0014faa0d 05020301 0fcc 5ee9 | KWD |
| 08/28/07 | TEST IN SYSTEM – PASS AWAITING FRONT METAL |  | KWD |
| 09/12/07 | FRONT METAL INSTALLED ON BOARD IN CABINET |  | KWD |
| 07/22/09 | Implemented ECO | MNSN-0119 | KZ |
| 07/23/09 | Updated Firmware: CLKBIASFPGAv44.mcs CLKBIASFPGAv44.bit CLKBIASCFGCPLDv42.jed CLKBIASSEQCPLDv401.jed | 0014f873c 08082104 0fcc 5ee9 | JRG |
| 07/23/09 | Retested Board, board passes functional test. | PASS | JRG |
| 8/06/09 | Installed into 1.5m Echelle system |  | R.G. |