## Detector Signals
E2V CCD 231-84-0-E08 4K x 4K backside illuminated
|  |  | Serial Number | 08442-22-01
Connector 1
Pin | Signal | Description | Min | Typical | Tested | Max | Absolute Max relative to Vss
1 | SS | Substrate (see note 12) | 0 | 0 | 0 | 10 | N/A
2 | DOS-E | Dummy Output Source (E) | (see note 8) |  |  |  | N/A
3 | OS-E | Output Source (E) | (see note 8) |  |  |  | N/A
4 | OG-E | Output Gate (E) (see note 9) | 1 | 2 | 2/18 | (see note 9) | ±20
5 | DG-A | Dump Gate (A) (see note 11) | -2 | 0 | 0 | 0.5 | ±20
6 | RG-E | Reset Gate (E) (see note 13) | 9 | 12 | 10 | 14 | ±20
7 | SW-E | Summing Well (E) (see note 9) | 9 | 10 | 10/2 | 12 | ±20
8 | E1 | Register Clock Phase 1 (E) | 9 | 10 | 12 | 12 | ±20
9 | E2 | Register Clock Phase 2 (E) | 9 | 10 | 12 | 12 | ±20
10 | E3-F3 | Register Clock Phase 3 (E and F) | 9 | 10 | 12 | 12 | ±20
11 | F1 | Register Clock Phase 1 (F) | 9 | 10 | 12 | 12 | ±20
12 | F2 | Register Clock Phase 2 (F) | 9 | 10 | 12 | 12 | ±20
13 | SW-F | Summing Well (F) (see note 9) | 9 | 10 | 10/2 | 12 | ±20
14 | RG-F | Reset Gate (F) (see note 13) | 9 | 12 | 10 | 14 | ±20
15 | TG-A | Transfer Gate (A) | 9 | 10 | 10 | 12 | ±20
16 | OG-F | Output Gate (F) (see note 9) | 1 | 2 | 2/18 | (see note 9) | ±20
17 | OS-F | Output Source (F) | (see note 8) |  |  |  | N/A
18 | DOS-F | Dummy Output Source (F) | (see note 8) |  |  |  | N/A
19 | SS | Substrate (see note 12) | 0 | 0 | 0 | 10 | N/A
20 | DOD-E | Dummy Output Drain (E) | 27 | 29 | 27.5 | 31 | -0.3 TO +35
21 | RD-E | Reset Drain (E) | 16 | 17 | 17.0 | 19 | -0.3 TO +25
22 | OD-E | Output Drain (E) | 27 | 29 | 27.5 | 31 | -0.3 TO +35
23 | SS | Substrate (see note 12) | 0 | 0 | 0 | 10 | N/A
24 | A4 | Image Area Clock Phase 4 (A) | 9 | 10 | 10 | 12 | ±20
25 | A3 | Image Area Clock Phase 3 (A) | 9 | 10 | 10 | 12 | ±20
26 | B4 | Image Area Clock Phase 4 (B) | 9 | 10 | 10 | 12 | ±20
27 | B3 | Image Area Clock Phase 3 (B) | 9 | 10 | 10 | 12 | ±20
28 | SS | Substrate (see note 12) | 0 | 0 | 0 | 10 | N/A
29 | DD-A | Dump Drain (A) | 27 | 29 | 29 | 31 | -0.3 TO +35
30 | B1 | Image Area Clock Phase 1 (B) | 9 | 10 | 10 | 12 | ±20
31 | B2 | Image Area Clock Phase 2 (B) | 9 | 10 | 10 | 12 | ±20
32 | A1 | Image Area Clock Phase 1 (A) | 9 | 10 | 10 | 12 | ±20
33 | A2 | Image Area Clock Phase 2 (A) | 9 | 10 | 10 | 12 | ±20
34 | SS | Substrate (see note 12) | 0 | 0 | 0 | 10 | N/A
35 | OD-F | Output Drain (F) | 27 | 29 | 27.5 | 31 | -0.3 TO +35
36 | RD-F | Reset Drain (F) | 16 | 17 | 17.0 | 19 | -0.3 TO +25
37 | DOD-F | Dummy Output Drain (F) | 27 | 29 | 27.5 | 31 | -0.3 TO +35
Connector 2
Pin | Signal | Description | Min | Typical | Tested | Max | Absolute Max relative to Vss
1 | SS | Substrate (see note 12) | 0 | 0 | 0 | 10 | N/A
2 | DOS-G | Dummy Output Source (G) | (see note 8) |  |  |  | N/A
3 | OS-G | Output Source (G) | (see note 8) |  |  |  | N/A
4 | OG-G | Output Gate (G) (see note 9) | 1 | 2 | 2/18 | (see note 9) | ±20
5 | DG-D | Dump Gate (D) (see note 11) | -2 | 0 | 0 | 0.5 | ±20
6 | RG-G | Reset Gate (G) (see note 13) | 9 | 12 | 10 | 14 | ±20
7 | SW-G | Summing Well (G) (see note 9) | 9 | 10 | 10/2 | 12 | ±20
8 | G1 | Register Clock Phase 1 (G) | 9 | 10 | 12 | 12 | ±20
9 | G2 | Register Clock Phase 2 (G) | 9 | 10 | 12 | 12 | ±20
10 | G3-H3 | Register Clock Phase 3 (G and H) | 9 | 10 | 12 | 12 | ±20
11 | H1 | Register Clock Phase 1 (H) | 9 | 10 | 12 | 12 | ±20
12 | H2 | Register Clock Phase 2 (H) | 9 | 10 | 12 | 12 | ±20
13 | SW-H | Summing Well (H) (see note 9) | 9 | 10 | 10/2 | 12 | ±20
14 | RG-H | Reset Gate (H) (see note 13) | 9 | 12 | 10 | 14 | ±20
15 | TG-D | Transfer Gate (D) | 9 | 10 | 10 | 12 | ±20
16 | OG-H | Output Gate (H) (see note 9) | 1 | 2 | 2/18 | (see note 9) | ±20
17 | OS-H | Output Source (H) | (see note 8) |  |  |  | N/A
18 | DOS-H | Dummy Output Source (H) | (see note 8) |  |  |  | N/A
19 | SS | Substrate (see note 12) | 0 | 0 | 0 | 10 | N/A
20 | DOD-G | Dummy Output Drain (G) | 27 | 29 | 27.5 | 31 | -0.3 TO +35
21 | RD-G | Reset Drain (G) | 16 | 17 | 17.0 | 19 | -0.3 TO +25
22 | OD-G | Output Drain (G) | 27 | 29 | 27.5 | 31 | -0.3 TO +35
23 | SS | Substrate (see note 12) | 0 | 0 | 0 | 10 | N/A
24 | D1 | Image Area Clock Phase 1 (D) | 9 | 10 | 10 | 12 | ±20
25 | D2 | Image Area Clock Phase 2 (D) | 9 | 10 | 10 | 12 | ±20
26 | C1 | Image Area Clock Phase 1 (C) | 9 | 10 | 10 | 12 | ±20
27 | C2 | Image Area Clock Phase 2 (C) | 9 | 10 | 10 | 12 | ±20
28 | SS | Substrate (see note 12) | 0 | 0 | 10 | 10 | N/A
29 | DD-D | Dump Drain (D) | 27 | 29 | 29 | 31 | -0.3 TO +35
30 | C4 | Image Area Clock Phase 4 (C) | 9 | 10 | 10 | 12 | ±20
31 | C3 | Image Area Clock Phase 3 (C) | 9 | 10 | 10 | 12 | ±20
32 | D4 | Image Area Clock Phase 4 (D) | 9 | 10 | 10 | 12 | ±20
33 | D3 | Image Area Clock Phase 3 (D) | 9 | 10 | 10 | 12 | ±20
34 | SS | Substrate (see note 12) | 0 | 0 | 0 | 10 | N/A
35 | OD-H | Output Drain (H) | 27 | 29 | 27.5 | 31 | -0.3 TO +35
36 | RD-H | Reset Drain (H) | 16 | 17 | 17.0 | 19 | -0.3 TO +25
37 | DOD-H | Dummy Output Drain (H) | 27 | 29 | 27.5 | 31 | -0.3 TO +35
Notes:
8 | Do not connect to voltage supply but use an approx. 5 mA current source or an approx. 5 kΩ external load. The quiescent voltage on OS is then about 6 - 8 V above the reset drain voltage and is typically 24 V. The current through these pins must not exceed 20mA.
Permanent damage may result if, in operation, OS or DOS experience short circuit conditions.
For highest speed operation the output load resistor can be reduced from 5 kΩ to approximately 2.2 kΩ, but note that this will increase power consumption.
If the device is to be operated with a register clock period of below about 1 MHz then the load may be increased to 10 kΩ to reduce power consumption.
9 | Default operation (mode 1) shown with OG set to OG-Lo, with a +2 V nominal value. In this mode SW may be clocked as R∅3 if a summing well function is not required. OG-Lo should have a maximum value of +5V.
For alternative operation in a low responsivity mode (mode 2) with increased charge handling, OG should be set to OG-Hi and SW should be operated as OG-Lo (i.e. 2V typical). See below for appropriate OG-Hi values. Charge is now read out as R∅2 goes low.
See note 12 also for discussion about Substrate voltage (Vss). With high substrate voltage OG-Hi may be set to a nominal +20V, which offers best linearity in mode-2. With low substrate voltage the allowed maximum value of OG-Hi is limited to a nominal +18V; the lower OG-Hi value has a greater non-linearity.
10 | To ensure that any device can be operated the camera should be designed so that any value in the range “min” to “max” can be provided. All operating voltages are with respect to image clock low (nominally 0 V).
The clock pulse low levels should be in the range 0 ± 0.5 V for image clocks. The register and SW clock low level should be +1 V higher. Reset clock low may be nominally 0 V or +1 V.
In all cases, specific recommended settings will be supplied with each science-grade sensor.
11 | Non-charge dumping level shown. For charge dumping, DG should be pulsed to 12 ± 2 V.
12 | The substrate voltage (Vss) has a default recommended value of 0V (“low” substrate). This is particularly recommended for deep-depletion device variants, since it optimises depletion depth for best Point Spread Function. Devices may alternatively be operated at “high” substrate, with Vss=9V. The high substrate setting offers slightly lower dark current, although this is usually not of primary concern when the device is cryogenically cooled.
The substrate setting has some consequence for the allowed OG upper voltage level, as discussed in note 9.
13 | Standard silicon variants are expected to be used with ∅R at +10V or more; deep depletion variants require at least +12V. A higher value will give a correspondingly higher reset feedthrough signal in the device output (OS).
PIN CONNECTIONS (View facing pins of connector)
This numbering applies to all connectors. The connector is a Glenair 37P micro D type.

## Cryostat Connectors
Glenair |  |  | Hermetic
MWDM2L-37SSP |  |  | PT02H-20-41P
CCD Connector P1 |  |  | Cryostat Connector J1
Pin | Signal |  | Signal | Pin
1 | SS | 0.005" Cu. Twisted with J1:r | VSS | s
2 | DOS-E | Not Wired
3 | OS-E | 0.005" Cu. Twisted with J1:m | OS_E | n
4 | OG-E | 0.005" Const. | OG_E | R
5 | DG-A | 0.005" Const. | DG_A | T
6 | RG-E | 0.005" Const. | RG_E | i
7 | SW-E | 0.005" Const. | SW_E | X
8 | E1 | 0.005" Const. | E1 | V
9 | E2 | 0.005" Const. | E2 | W
10 | E3-F3 | 0.005" Const. | E3_F3 | A
11 | F1 | 0.005" Const. | F1 | B
12 | F2 | 0.005" Const. | F2 | C
13 | SW-F | 0.005" Const. | SW_F | Y
14 | RG-F | 0.005" Const. | RG_F | j
15 | TG-A | 0.005" Const. | TG_A | Z
16 | OG-F | 0.005" Const. | OG_F | S
17 | OS-F | 0.005" Cu. Twisted with J1:s | OS_F | r
18 | DOS-F | Not Wired
19 | SS | 0.005" Cu. Twisted withJ1:n | VSS | m
20 | DOD-E | Not Wired |  | f
21 | RD-E | 0.005" Const. | RD_E | N
22 | OD-E | 0.010" Const. | OD_E | p
23 | SS | 0.005" Cu. In parallel with P1:34 | VSS | t
24 | A4 | 0.010" Const. In parallel with P1:26 | A4 | G
25 | A3 | 0.010" Const. In parallel with P1:27 | A3 | F
26 | B4 | 0.010" Const. In parallel with P1:24 | B4 | G
27 | B3 | 0.010" Const. In parallel with P1:25 | B3 | F
28 | SS | 0.005" Cu. | VSS | k
29 | DD-A | 0.005" Const. | DD_A | U
30 | B1 | 0.010" Const. In parallel with P1:32 | B1 | D
31 | B2 | 0.010" Const. In parallel with P1:33 | B2 | E
32 | A1 | 0.010" Const. In parallel with P1:30 | A1 | D
33 | A2 | 0.010" Const. In parallel with P1:31 | A2 | E
34 | SS | 0.005" Cu. In parallel with P1:23 | VSS | t
35 | OD-F | 0.010" Const. | OD_F | q
36 | RD-F | 0.005" Const. | RD_F | P
37 | DOD-F | Not Wired |  | g
|  | Not Wired | SHLD | M
|  | Not Wired | VID_SHLD | h
|  |  | TS1P | H
|  |  | TS1N | J
|  |  | HTR_P | K
|  |  | HTR_N | L
|  |  | TS2P | d
|  |  | TS2N | e
Glenair |  |  | Hermetic
MWDM2L-37SSP |  |  | PT02H-20-41P
CCD Connector P2 |  |  | Cryostat Connector J2
Pin | Signal |  | Signal | Pin
1 | SS | 0.005" Cu. Twisted with J2:r | VSS | s
2 | DOS-G | Not Wired
3 | OS-G | 0.005" Cu. Twisted with J2:s | OS_G | r
4 | OG-G | 0.005" Const. | OG_G | S
5 | DG-D | 0.005" Const. | DG_G | T
6 | RG-G | 0.005" Const. | RG_G | j
7 | SW-G | 0.005" Const. | SW_G | Y
8 | G1 | 0.005" Const. | G1 | B
9 | G2 | 0.005" Const. | G2 | C
10 | G3-H3 | 0.005" Const. | H3_G3 | A
11 | H1 | 0.005" Const. | H1 | V
12 | H2 | 0.005" Const. | H2 | W
13 | SW-H | 0.005" Const. | SW_H | X
14 | RG-H | 0.005" Const. | RG_H | i
15 | TG-D | 0.005" Const. | TG_D | Z
16 | OG-H | 0.005" Const. | OG_H | R
17 | OS-H | 0.005" Cu. Twisted with J2:m | OS_H | n
18 | DOS-H | Not Wired
19 | SS | 0.005" Cu. Twisted with J2:n | VSS | m
20 | DOD-G | Not Wired |  | g
21 | RD-G | 0.005" Const. | RD_G | P
22 | OD-G | 0.010" Const. | OD_G | q
23 | SS | 0.005" Cu. In parallel with P2:34 | VSS | t
24 | D1 | 0.010" Const. In parallel with P2:26 | D1 | D
25 | D2 | 0.010" Const. In parallel with P2:27 | D2 | E
26 | C1 | 0.010" Const. In parallel with P2:24 | C1 | D
27 | C2 | 0.010" Const. In parallel with P2:25 | C2 | E
28 | SS | 0.005" Cu. | VSS | k
29 | DD-D | 0.005" Const. | DD_D | U
30 | C4 | 0.010" Const. In parallel with P2:32 | C4 | G
31 | C3 | 0.010" Const. In parallel with P2:33 | C3 | F
32 | D4 | 0.010" Const. In parallel with P2:30 | D4 | G
33 | D3 | 0.010" Const. In parallel with P2:31 | D3 | F
34 | SS | 0.005" Cu. In parallel with P2:23 | VSS | t
35 | OD-H | 0.010" Const. | OD_H | p
36 | RD-H | 0.005" Const. | RD_H | N
37 | DOD-H | Not Wired |  | f
|  | Not Wired | SHLD | M
Ainsley Htr Conn |  | Not Wired | VID_SHLD | h
1 | D1_Anode | 0.010" Const. | TS1P | H
2 | D1_Cath | 0.010" Const. | TS1N | J
7 | Fuse | 0.010" Const. | HTR_P | K
8 | R1 | 0.010" Const. | HTR_N | L
3 | D2_Anode | 0.010" Const. | TS2P | d
4 | D2_Cath | 0.010" Const. | TS2N | e

## Detector Cables
Amphenol |  |  |  | Amphenol
PT06A-20-41S-SR |  |  |  | PT06A-20-41P-SR
Cryostat Connector P1 |  |  |  | Chassis Connector P3
Pin | Signal | Wire Type | Comments | Signal | Pin
A | E3_F3 | 22awg (19/32) Tfe |  | E3_F3 | A
B | F1 | 24awg (19/36) Tfe |  | F1 | B
C | F2 | 24awg (19/36) Tfe |  | F2 | C
D | A1_B1 | 24awg (19/36) Tfe |  | A1_B1 | D
E | A2_B2 | 24awg (19/36) Tfe |  | A2_B2 | E
F | A3_B3 | 24awg (19/36) Tfe |  | A3_B3 | F
G | A4_B4 | 24awg (19/36) Tfe |  | A4_B4 | G
H | TS1P | 24awg (19/36) Tfe |  | TS1P | H
J | TS1N | 24awg (19/36) Tfe |  | TS1N | J
K | HTR_P | 20Awg (19/32) Pvc |  | HTR_P | K
L | HTR_N | 20Awg (19/32) Pvc |  | HTR_N | L
M |  |  | All shields tied to this point | SHLD | M
N | RD_E | 22awg (19/32) Tfe |  | RD_E | N
P | RD_F | 22awg (19/32) Tfe |  | RD_F | P
R | OG_E | 22awg (19/32) Tfe |  | OG_E | R
S | OG_F | 22awg (19/32) Tfe |  | OG_F | S
T | DG_A | 24awg (19/36) Tfe |  | DG_A | T
U | DD_A | 22awg (19/32) Tfe |  | DD_A | U
V | E1 | 24awg (19/36) Tfe |  | E1 | V
W | E2 | 24awg (19/36) Tfe |  | E2 | W
X | SW_E | 24awg (19/36) Tfe |  | SW_E | X
Y | SW_F | 24awg (19/36) Tfe |  | SW_F | Y
Z | TG_A | 24awg (19/36) Tfe |  | TG_A | Z
a |  |  |  |  | a
b |  |  |  |  | b
c |  |  |  |  | c
d | TS2P | 24awg (19/36) Tfe |  | TS2P | d
e | TS2N | 24awg (19/36) Tfe |  | TS2N | e
f | DOD_E | twisted pair 1 (Alpha 2461) | Red conductor | DOD_E | f
g | DOD_F | twisted pair 1 (Alpha 2461) | Black conductor | DOD_F | g
h |  |  | All twisted pair shields to this point | VID_SHLD | h
i | RG_E | 24awg (19/36) Tfe |  | RG_E | i
j | RG_F | 24awg (19/36) Tfe |  | RG_F | j
k | Vss | 20Awg (19/32) Pvc |  | Vss | k
m | Vss | twisted pair 2 (Alpha 2461) | Black conductor | Vss | m
n | OS_E | twisted pair 2 (Alpha 2461) | Red conductor | OS_E | n
p | OD_E | twisted pair 3 (Alpha 2461) | Red conductor | OD_E | p
q | OD_F | twisted pair 3 (Alpha 2461) | Black conductor | OD_F | q
r | OS_F | twisted pair 4 (Alpha 2461) | Red conductor | OS_F | r
s | Vss | twisted pair 4 (Alpha 2461) | Black conductor | Vss | s
t | Vss | 20Awg (19/32) Pvc |  | Vss | t
Amphenol |  |  |  | Amphenol
PT06A-20-41S-SR |  |  |  | PT06A-20-41P-SR
Cryostat Connector P2 |  |  |  | Chassis ConnectorP4
Pin | Signal | Wire Type | Comments | Signal | Pin
A | H3_G3 | 22awg (19/32) Tfe |  | H3_G3 | A
B | G1 | 24awg (19/36) Tfe |  | G1 | B
C | G2 | 24awg (19/36) Tfe |  | G2 | C
D | C1_D1 | 24awg (19/36) Tfe |  | C1_D1 | D
E | C2_D2 | 24awg (19/36) Tfe |  | C2_D2 | E
F | C3_D3 | 24awg (19/36) Tfe |  | C3_D3 | F
G | C4_D4 | 24awg (19/36) Tfe |  | C4_D4 | G
H | TS1P | 24awg (19/36) Tfe | Twist with pin J | TS1P | H
J | TS1N | 24awg (19/36) Tfe | Twist with pin H | TS1N | J
K | HTR_P | 20Awg (19/32) Pvc | Twist with pin L | HTR_P | K
L | HTR_N | 20Awg (19/32) Pvc | Twist with pin K | HTR_N | L
M |  |  | All shields tied to this point | SHLD | M
N | RD_H | 22awg (19/32) Tfe |  | RD_H | N
P | RD_G | 22awg (19/32) Tfe |  | RD_G | P
R | OG_H | 22awg (19/32) Tfe |  | OG_H | R
S | OG_G | 22awg (19/32) Tfe |  | OG_G | S
T | DG_D | 24awg (19/36) Tfe |  | DG_D | T
U | DD_D | 22awg (19/32) Tfe |  | DD_D | U
V | H1 | 24awg (19/36) Tfe |  | H1 | V
W | H2 | 24awg (19/36) Tfe |  | H2 | W
X | SW_H | 24awg (19/36) Tfe |  | SW_H | X
Y | SW_G | 24awg (19/36) Tfe |  | SW_G | Y
Z | TG_D | 24awg (19/36) Tfe |  | TG_D | Z
a |  |  |  |  | a
b |  |  |  |  | b
c |  |  |  |  | c
d | TS2P | 24awg (19/36) Tfe | Twist with pin e | TS2P | d
e | TS2N | 24awg (19/36) Tfe | Twist with pin d | TS2N | e
f | DOD_H | twisted pair 1 (Alpha 2461) | Red conductor | DOD_H | f
g | DOD_G | twisted pair 1 (Alpha 2461) | Black conductor | DOD_G | g
h |  |  | All twisted pair shields to this point | VID_SHLD | h
i | RG_H | 24awg (19/36) Tfe |  | RG_H | i
j | RG_G | 24awg (19/36) Tfe |  | RG_G | j
k | Vss | 20Awg (19/32) Pvc |  | Vss | k
m | Vss | twisted pair 1 (Alpha 2461) | Black conductor | Vss | m
n | OS_H | twisted pair 1 (Alpha 2461) | Red conductor | OS_H | n
p | OD_H | twisted pair 2 (Alpha 2461) | Red conductor | OD_H | p
q | OD_G | twisted pair 2 (Alpha 2461) | Black conductor | OD_G | q
r | OS_G | twisted pair 3 (Alpha 2461) | Red conductor | OS_G | r
s | VSS | twisted pair 3 (Alpha 2461) | Black conductor | VSS | s
t | VSS | 20Awg (19/32) Pvc |  | VSS | t

## Chassis Connectors
PT02-20-41S
Chassis Connector J3 |  |  | MONSOON Connectors
Pin | Signal | Signal | Slot | Connector | Pin | Hardware Signal
A | E3_F3 | S3ef | 3 | P2 | 10 | CLKOUT24
B | F1 | S1f | 3 | P2 | 13 | CLKOUT21
C | F2 | S2f | 3 | P2 | 11 | CLKOUT23
D | A1_B1 | P1ab | 3 | P2 | 6 | CLKOUT28
E | A2_B2 | P2ab | 3 | P2 | 5 | CLKOUT29
F | A3_B3 | P3ab | 3 | P2 | 4 | CLKOUT30
G | A4_B4 | P4ab | 3 | P2 | 3 | CLKOUT31
H
J
K
L
M | SHLD
N | RD_E | RDe | 2 | P3 | 14 | HVBIAS20
P | RD_F | RDf | 2 | P3 | 12 | HVBIAS22
R | OG_E | OGe | 2 | P3 | 26 | HVBIAS12
S | OG_F | OGf | 2 | P3 | 24 | HVBIAS14
T | DG_A | DGab | 3 | P2 | 8 | CLKOUT26
U | DD_A | DDab | 2 | P3 | 8 | HVBIAS26
V | E1 | S1e | 3 | P2 | 14 | CLKOUT20
W | E2 | S2e | 3 | P2 | 12 | CLKOUT22
X | SW_E | SWe | 3 | P2 | 16 | CLKOUT18
Y | SW_F | SWf | 3 | P2 | 15 | CLKOUT19
Z | TG_A | TGab | 3 | P2 | 9 | CLKOUT25
a
b
c
d
e
f | DOD_E
g | DOD_F
h | VID_SHLD
i | RG_E | RGe | 3 | P2 | 18 | CLKOUT16
j | RG_F | RGf | 3 | P2 | 17 | CLKOUT17
k | VSS | GND | 2 | P3 | 21 | HVBIAS_RTN4
m | VSS | RSe | 2 | P4 | 14 | CHAN5-
n | OS_E | OSe | 2 | P4 | 13 | CHAN5+
p | OD_E | ODe | 2 | P3 | 34 | HVBIAS04
q | OD_F | ODf | 2 | P3 | 32 | HVBIAS06
r | OS_F | OSf | 2 | P4 | 17 | CHAN7+
s | VSS | RSf | 2 | P4 | 18 | CHAN7-
t | VSS | GND | 3 | P2 | 39 | CLK_RTN6
PT02A-20-41S
Chassis Connector J4 |  |  | MONSOON Connectors
Pin | Signal | Signal | Slot | Connector | Pin
A | G3_H3 | S3gh | 3 | P2 | 30 | CLKOUT08
B | G1 | S1g | 3 | P2 | 33 | CLKOUT05
C | G2 | S2g | 3 | P2 | 31 | CLKOUT07
D | C1_D1 | P1cd | 3 | P2 | 26 | CLKOUT12
E | C2_D2 | P2cd | 3 | P2 | 25 | CLKOUT13
F | C3_D3 | P3cd | 3 | P2 | 24 | CLKOUT14
G | C4_D4 | P4cd | 3 | P2 | 23 | CLKOUT15
H | TS1P
J | TS1N
K | HTR_P
L | HTR_N
M | SHLD
N | RD_H | RDh | 2 | P3 | 18 | HVBIAS16
P | RD_G | RDg | 2 | P3 | 16 | HVBIAS18
R | OG_H | OGh | 2 | P3 | 30 | HVBIAS08
S | OG_G | OGg | 2 | P3 | 28 | HVBIAS10
T | DG_G | DGcd | 3 | P2 | 28 | CLKOUT10
U | DD_D | DDcd | 2 | P3 | 10 | HVBIAS24
V | H1 | S1h | 3 | P2 | 34 | CLKOUT04
W | H2 | S2h | 3 | P2 | 32 | CLKOUT06
X | SW_H | SWh | 3 | P2 | 36 | CLKOUT02
Y | SW_G | SWg | 3 | P2 | 35 | CLKOUT03
Z | TG_D | TGcd | 3 | P2 | 29 | CLKOUT09
a
b
c
d | TS2P
e | TS2N
f | DOD_H
g | DOD_G
h | VID_SHLD
i | RG_H | RGh | 3 | P2 | 38 | CLKOUT00
j | RG_G | RGg | 3 | P2 | 37 | CLKOUT01
k | VSS | GND | 2 | P3 | 20 | HVBIAS_RTN3
m | VSS | RSh | 2 | P4 | 6 | CHAN1-
n | OS_H | OSh | 2 | P4 | 5 | CHAN1+
p | OD_H | ODh | 2 | P3 | 38 | HVBIAS00
q | OD_G | ODg | 2 | P3 | 36 | HVBIAS02
r | OS_G | OSg | 2 | P4 | 9 | CHAN3+
s | VSS | RSg | 2 | P4 | 10 | CHAN3-
t | VSS | GND | 3 | P2 | 40 | CLK_RTN7

## MONSOON Orange Signals
MONSOON Board CLK & Bias
Slot | Connector | Pin | Board | Hardware Signal | Function Name
3 | P2 | 1 | CLK & BIAS | CLK_RTN0
3 | P2 | 2 | CLK & BIAS | CLK_RTN1
3 | P2 | 3 | CLK & BIAS | CLKOUT31 | CLK_OUT_31
3 | P2 | 4 | CLK & BIAS | CLKOUT30 | CLK_OUT_30
3 | P2 | 5 | CLK & BIAS | CLKOUT29 | CLK_OUT_29
3 | P2 | 6 | CLK & BIAS | CLKOUT28 | CLK_OUT_28
3 | P2 | 7 | CLK & BIAS | CLKOUT27 | CLK_OUT_27
3 | P2 | 8 | CLK & BIAS | CLKOUT26 | CLK_OUT_26
3 | P2 | 9 | CLK & BIAS | CLKOUT25 | CLK_OUT_25
3 | P2 | 10 | CLK & BIAS | CLKOUT24 | CLK_OUT_24
3 | P2 | 11 | CLK & BIAS | CLKOUT23 | CLK_OUT_23
3 | P2 | 12 | CLK & BIAS | CLKOUT22 | CLK_OUT_22
3 | P2 | 13 | CLK & BIAS | CLKOUT21 | CLK_OUT_21
3 | P2 | 14 | CLK & BIAS | CLKOUT20 | CLK_OUT_20
3 | P2 | 15 | CLK & BIAS | CLKOUT19 | CLK_OUT_19
3 | P2 | 16 | CLK & BIAS | CLKOUT18 | CLK_OUT_18
3 | P2 | 17 | CLK & BIAS | CLKOUT17 | CLK_OUT_17
3 | P2 | 18 | CLK & BIAS | CLKOUT16 | CLK_OUT_16
3 | P2 | 19 | CLK & BIAS | CLK_RTN2
3 | P2 | 20 | CLK & BIAS | CLK_RTN3
3 | P2 | 21 | CLK & BIAS | CLK_RTN4
3 | P2 | 22 | CLK & BIAS | CLK_RTN5
3 | P2 | 23 | CLK & BIAS | CLKOUT15 | CLK_OUT_15
3 | P2 | 24 | CLK & BIAS | CLKOUT14 | CLK_OUT_14
3 | P2 | 25 | CLK & BIAS | CLKOUT13 | CLK_OUT_13
3 | P2 | 26 | CLK & BIAS | CLKOUT12 | CLK_OUT_12
3 | P2 | 27 | CLK & BIAS | CLKOUT11 | CLK_OUT_11
3 | P2 | 28 | CLK & BIAS | CLKOUT10 | CLK_OUT_10
3 | P2 | 29 | CLK & BIAS | CLKOUT09 | CLK_OUT_09
3 | P2 | 30 | CLK & BIAS | CLKOUT08 | CLK_OUT_08
3 | P2 | 31 | CLK & BIAS | CLKOUT07 | CLK_OUT_07
3 | P2 | 32 | CLK & BIAS | CLKOUT06 | CLK_OUT_06
3 | P2 | 33 | CLK & BIAS | CLKOUT05 | CLK_OUT_05
3 | P2 | 34 | CLK & BIAS | CLKOUT04 | CLK_OUT_04
3 | P2 | 35 | CLK & BIAS | CLKOUT03 | CLK_OUT_03
3 | P2 | 36 | CLK & BIAS | CLKOUT02 | CLK_OUT_02
3 | P2 | 37 | CLK & BIAS | CLKOUT01 | CLK_OUT_01
3 | P2 | 38 | CLK & BIAS | CLKOUT00 | CLK_OUT_00
3 | P2 | 39 | CLK & BIAS | CLK_RTN6
3 | P2 | 40 | CLK & BIAS | CLK_RTN7
3 | P3 | 1 | CLK & BIAS | LVBIAS_RTN0
3 | P3 | 2 | CLK & BIAS | LVBIAS_RTN1
3 | P3 | 3 | CLK & BIAS | LVBIAS35 | CLK_BIAS[35]
3 | P3 | 4 | CLK & BIAS | LVBIAS34 | CLK_BIAS[34]
3 | P3 | 5 | CLK & BIAS | LVBIAS33 | CLK_BIAS[33]
3 | P3 | 6 | CLK & BIAS | LVBIAS32 | CLK_BIAS[32]
3 | P3 | 7 | CLK & BIAS | LVBIAS31 | CLK_BIAS[31]
3 | P3 | 8 | CLK & BIAS | LVBIAS30 | CLK_BIAS[30]
3 | P3 | 9 | CLK & BIAS | LVBIAS29 | CLK_BIAS[29]
3 | P3 | 10 | CLK & BIAS | LVBIAS28 | CLK_BIAS[28]
3 | P3 | 11 | CLK & BIAS | LVBIAS27 | CLK_BIAS[27]
3 | P3 | 12 | CLK & BIAS | LVBIAS26 | CLK_BIAS[26]
3 | P3 | 13 | CLK & BIAS | LVBIAS25 | CLK_BIAS[25]
3 | P3 | 14 | CLK & BIAS | LVBIAS24 | CLK_BIAS[24]
3 | P3 | 15 | CLK & BIAS | LVBIAS23 | CLK_BIAS[23]
3 | P3 | 16 | CLK & BIAS | LVBIAS22 | CLK_BIAS[22]
3 | P3 | 17 | CLK & BIAS | LVBIAS21 | CLK_BIAS[21]
3 | P3 | 18 | CLK & BIAS | LVBIAS20 | CLK_BIAS[20]
3 | P3 | 19 | CLK & BIAS | LVBIAS19 | CLK_BIAS[19]
3 | P3 | 20 | CLK & BIAS | LVBIAS18 | CLK_BIAS[18]
3 | P3 | 21 | CLK & BIAS | LVBIAS17 | CLK_BIAS[17]
3 | P3 | 22 | CLK & BIAS | LVBIAS16 | CLK_BIAS[16]
3 | P3 | 23 | CLK & BIAS | LVBIAS15 | CLK_BIAS[15]
3 | P3 | 24 | CLK & BIAS | LVBIAS14 | CLK_BIAS[14]
3 | P3 | 25 | CLK & BIAS | LVBIAS13 | CLK_BIAS[13]
3 | P3 | 26 | CLK & BIAS | LVBIAS12 | CLK_BIAS[12]
3 | P3 | 27 | CLK & BIAS | LVBIAS11 | CLK_BIAS[11]
3 | P3 | 28 | CLK & BIAS | LVBIAS10 | CLK_BIAS[10]
3 | P3 | 29 | CLK & BIAS | LVBIAS09 | CLK_BIAS[9]
3 | P3 | 30 | CLK & BIAS | LVBIAS08 | CLK_BIAS[8]
3 | P3 | 31 | CLK & BIAS | LVBIAS07 | CLK_BIAS[7]
3 | P3 | 32 | CLK & BIAS | LVBIAS06 | CLK_BIAS[6]
3 | P3 | 33 | CLK & BIAS | LVBIAS05 | CLK_BIAS[5]
3 | P3 | 34 | CLK & BIAS | LVBIAS04 | CLK_BIAS[4]
3 | P3 | 35 | CLK & BIAS | LVBIAS03 | CLK_BIAS[3]
3 | P3 | 36 | CLK & BIAS | LVBIAS02 | CLK_BIAS[2]
3 | P3 | 37 | CLK & BIAS | LVBIAS01 | CLK_BIAS[1]
3 | P3 | 38 | CLK & BIAS | LVBIAS00 | CLK_BIAS[0]
3 | P3 | 39 | CLK & BIAS | LVBIAS_RTN2
3 | P3 | 40 | CLK & BIAS | LVBIAS_RTN3
3 | P3 | 41 | CLK & BIAS | FBIAS7 | CLK_FBIAS[7]
3 | P3 | 42 | CLK & BIAS | FBIAS6 | CLK_FBIAS[6]
3 | P3 | 43 | CLK & BIAS | FBIAS5 | CLK_FBIAS[5]
3 | P3 | 44 | CLK & BIAS | FBIAS4 | CLK_FBIAS[4]
3 | P3 | 45 | CLK & BIAS | FBIAS3 | CLK_FBIAS[3]
3 | P3 | 46 | CLK & BIAS | FBIAS2 | CLK_FBIAS[2]
3 | P3 | 47 | CLK & BIAS | FBIAS1 | CLK_FBIAS[1]
3 | P3 | 48 | CLK & BIAS | FBIAS0 | CLK_FBIAS[0]
3 | P3 | 49 | CLK & BIAS | FBIAS_RTN0
3 | P3 | 50 | CLK & BIAS | FBIAS_RTN1
2 | P2 | 1 | CCD ACQ | DOP_RTN0
2 | P2 | 2 | CCD ACQ | DOP_RTN1
2 | P2 | 3 | CCD ACQ | DOP31 | CCD_DOPREG[31]
2 | P2 | 4 | CCD ACQ | DOP30 | CCD_DOPREG[30]
2 | P2 | 5 | CCD ACQ | DOP29 | CCD_DOPREG[29]
2 | P2 | 6 | CCD ACQ | DOP28 | CCD_DOPREG[28]
2 | P2 | 7 | CCD ACQ | DOP27 | CCD_DOPREG[27]
2 | P2 | 8 | CCD ACQ | DOP26 | CCD_DOPREG[26]
2 | P2 | 9 | CCD ACQ | DOP25 | CCD_DOPREG[25]
2 | P2 | 10 | CCD ACQ | DOP24 | CCD_DOPREG[24]
2 | P2 | 11 | CCD ACQ | DOP23 | CCD_DOPREG[23]
2 | P2 | 12 | CCD ACQ | DOP22 | CCD_DOPREG[22]
2 | P2 | 13 | CCD ACQ | DOP21 | CCD_DOPREG[21]
2 | P2 | 14 | CCD ACQ | DOP20 | CCD_DOPREG[20]
2 | P2 | 15 | CCD ACQ | DOP19 | CCD_DOPREG[19]
2 | P2 | 16 | CCD ACQ | DOP18 | CCD_DOPREG[18]
2 | P2 | 17 | CCD ACQ | DOP17 | CCD_DOPREG[17]
2 | P2 | 18 | CCD ACQ | DOP16 | CCD_DOPREG[16]
2 | P2 | 19 | CCD ACQ | DOP15 | CCD_DOPREG[15]
2 | P2 | 20 | CCD ACQ | DOP14 | CCD_DOPREG[14]
2 | P2 | 21 | CCD ACQ | DOP13 | CCD_DOPREG[13]
2 | P2 | 22 | CCD ACQ | DOP12 | CCD_DOPREG[12]
2 | P2 | 23 | CCD ACQ | DOP11 | CCD_DOPREG[11]
2 | P2 | 24 | CCD ACQ | DOP10 | CCD_DOPREG[10]
2 | P2 | 25 | CCD ACQ | DOP09 | CCD_DOPREG[9]
2 | P2 | 26 | CCD ACQ | DOP08 | CCD_DOPREG[8]
2 | P2 | 27 | CCD ACQ | DOP07 | CCD_DOPREG[7]
2 | P2 | 28 | CCD ACQ | DOP06 | CCD_DOPREG[6]
2 | P2 | 29 | CCD ACQ | DOP05 | CCD_DOPREG[5]
2 | P2 | 30 | CCD ACQ | DOP04 | CCD_DOPREG[4]
2 | P2 | 31 | CCD ACQ | DOP03 | CCD_DOPREG[3]
2 | P2 | 32 | CCD ACQ | DOP02 | CCD_DOPREG[2]
2 | P2 | 33 | CCD ACQ | DOP01 | CCD_DOPREG[1]
2 | P2 | 34 | CCD ACQ | DOP00 | CCD_DOPREG[0]
2 | P3 | 1 | CCD ACQ | HVBIAS_RTN0
2 | P3 | 2 | CCD ACQ | HVBIAS_RTN1
2 | P3 | 3 | CCD ACQ | HVBIAS31 | CCD_BIAS[31]
2 | P3 | 4 | CCD ACQ | HVBIAS30 | CCD_BIAS[30]
2 | P3 | 5 | CCD ACQ | HVBIAS29 | CCD_BIAS[29]
2 | P3 | 6 | CCD ACQ | HVBIAS28 | CCD_BIAS[28]
2 | P3 | 7 | CCD ACQ | HVBIAS27 | CCD_BIAS[27]
2 | P3 | 8 | CCD ACQ | HVBIAS26 | CCD_BIAS[26]
2 | P3 | 9 | CCD ACQ | HVBIAS25 | CCD_BIAS[25]
2 | P3 | 10 | CCD ACQ | HVBIAS24 | CCD_BIAS[24]
2 | P3 | 11 | CCD ACQ | HVBIAS23 | CCD_BIAS[23]
2 | P3 | 12 | CCD ACQ | HVBIAS22 | CCD_BIAS[22]
2 | P3 | 13 | CCD ACQ | HVBIAS21 | CCD_BIAS[21]
2 | P3 | 14 | CCD ACQ | HVBIAS20 | CCD_BIAS[20]
2 | P3 | 15 | CCD ACQ | HVBIAS19 | CCD_BIAS[19]
2 | P3 | 16 | CCD ACQ | HVBIAS18 | CCD_BIAS[18]
2 | P3 | 17 | CCD ACQ | HVBIAS17 | CCD_BIAS[17]
2 | P3 | 18 | CCD ACQ | HVBIAS16 | CCD_BIAS[16]
2 | P3 | 19 | CCD ACQ | HVBIAS_RTN2
2 | P3 | 20 | CCD ACQ | HVBIAS_RTN3
2 | P3 | 21 | CCD ACQ | HVBIAS_RTN4
2 | P3 | 22 | CCD ACQ | HVBIAS_RTN5
2 | P3 | 23 | CCD ACQ | HVBIAS15 | CCD_BIAS[15]
2 | P3 | 24 | CCD ACQ | HVBIAS14 | CCD_BIAS[14]
2 | P3 | 25 | CCD ACQ | HVBIAS13 | CCD_BIAS[13]
2 | P3 | 26 | CCD ACQ | HVBIAS12 | CCD_BIAS[12]
2 | P3 | 27 | CCD ACQ | HVBIAS11 | CCD_BIAS[11]
2 | P3 | 28 | CCD ACQ | HVBIAS10 | CCD_BIAS[10]
2 | P3 | 29 | CCD ACQ | HVBIAS09 | CCD_BIAS[9]
2 | P3 | 30 | CCD ACQ | HVBIAS08 | CCD_BIAS[8]
2 | P3 | 31 | CCD ACQ | HVBIAS07 | CCD_BIAS[7]
2 | P3 | 32 | CCD ACQ | HVBIAS06 | CCD_BIAS[6]
2 | P3 | 33 | CCD ACQ | HVBIAS05 | CCD_BIAS[5]
2 | P3 | 34 | CCD ACQ | HVBIAS04 | CCD_BIAS[4]
2 | P3 | 35 | CCD ACQ | HVBIAS03 | CCD_BIAS[3]
2 | P3 | 36 | CCD ACQ | HVBIAS02 | CCD_BIAS[2]
2 | P3 | 37 | CCD ACQ | HVBIAS01 | CCD_BIAS[1]
2 | P3 | 38 | CCD ACQ | HVBIAS00 | CCD_BIAS[0]
2 | P3 | 39 | CCD ACQ | HVBIAS_RTN6
2 | P3 | 40 | CCD ACQ | HVBIAS_RTN7
2 | P4 | 1 | CCD ACQ | VID_RTN0
2 | P4 | 2 | CCD ACQ | VID_RTN1
2 | P4 | 3 | CCD ACQ | CHAN0+ | CCD_ADCDATA[0]
2 | P4 | 4 | CCD ACQ | CHAN0-
2 | P4 | 5 | CCD ACQ | CHAN1+ | CCD_ADCDATA[1]
2 | P4 | 6 | CCD ACQ | CHAN1-
2 | P4 | 7 | CCD ACQ | CHAN2+ | CCD_ADCDATA[2]
2 | P4 | 8 | CCD ACQ | CHAN2-
2 | P4 | 9 | CCD ACQ | CHAN3+ | CCD_ADCDATA[3]
2 | P4 | 10 | CCD ACQ | CHAN3-
2 | P4 | 11 | CCD ACQ | CHAN4+ | CCD_ADCDATA[4]
2 | P4 | 12 | CCD ACQ | CHAN4-
2 | P4 | 13 | CCD ACQ | CHAN5+ | CCD_ADCDATA[5]
2 | P4 | 14 | CCD ACQ | CHAN5-
2 | P4 | 15 | CCD ACQ | CHAN6+ | CCD_ADCDATA[6]
2 | P4 | 16 | CCD ACQ | CHAN6-
2 | P4 | 17 | CCD ACQ | CHAN7+ | CCD_ADCDATA[7]
2 | P4 | 18 | CCD ACQ | CHAN7-
2 | P4 | 19 | CCD ACQ | VID_RTN2
2 | P4 | 20 | CCD ACQ | VID_RTN3

## MONSOON Jumper Config
MONSOON Orange System Jumper and Option Settings
CCD Acquisition Board |  |  | SN 14
Bias Voltage Strapping
Bias Bank | Bias Channels | Min Voltage | Max Voltage | +Ve Jumper | -Ve Jumper | +Ve Resistor | -Ve Resistor
0 | 0 -> 7 | 0 | +25 | +VAB0 | -VAB0 | R634 | R671
1 | 8 - > 15 | -10 | +10 | +VAB1 | -VAB1 | R237 | R281
2 | 16 -> 23 | 0 | +25 | +VAB2 | -VAB2 | R240 | R274
3 | 24 -> 31 | 0 | +25 | +VAB3 | -VAB3 | R635 | R672
Bias DAC Reference Strapping
Signal |  | Bias Bank | Value | Jumper | Position | Set Resistor | Value
BIASDACREF0+ |  | 0 | +8.3 | JP9 | 2-3 | R740 | 2.32K | Ode,f,g,h
BIASDACREF0- |  | 0 | 0 | JP6 | 1-2 | R731 | 4.00K
BIASDACREF1+ |  | 1 | +4.12v | JP5 | 2-3 | R730 | 653 | Oge,f,g,h
BIASDACREF1- |  | 1 | -4.12v | JP10 | 2-3 | R742 | 1.65K
BIASDACREF2+ |  | 2 | +8.3v | JP8 | 2-3 | R270 | 2.32K | Rde,f,g,h
BIASDACREF2- |  | 2 | 0 | JP4 | 1-2 | R263 | 4.0K
BIASDACREF3+ |  | 3 | +8.3v | JP3 | 2-3 | R264 | 2.32K | Ddab,cd
BIASDACREF3- |  | 3 | 0 | JP7 | 1-2 | R271 | 4.0K
Video chain Voltage Strapping
Stage | Function | Min Voltage | Max Voltage | +Ve Jumper | -Ve Jumper | +Ve Resistor | -Ve Resistor
0 | Input Buffr | -15 | +5 | +W0 | -W0 | R170 | R168
1 | Inv Amp | +5 | -15 | +W1 | -W1 | R183 | R180
2 | Intergrator | +15 | -5 | +W2 | -W2 | R186 | R190
Video Offset Dac Range
Signal |  | Value | Jumper | Position
OFFDACREF+ |  | +5v | JP1 | 2-3
OFFDACREF- |  | -5v | JP2 | 2-3
Power Supply Selection
Signal |  | Value | Jumper | Position
+5VDIN Enable |  | Disable | JP14 | Open
+1.8V Src Select |  | External | JP12 | 1-2
Tst Osc Enable |  | Disable | JP13 | Open
Modifications to CCD Acq board for Chiron Application
Remove |  | Add |  |  | Purpose
C10 |  | Jumper TP1 to TP46 |  |  | Parallel Channel 0 and Channel 1 Detector Signals
C75 |  | Jumper TP8 to TP54 |  |  | Parallel Channel 2 and Channel 3 Detector Signals
C139 |  | Jumper TP16 to TP61 |  |  | Parallel Channel 4 and Channel 5 Detector Signals
C196 |  | Jumper TP24 to TP68 |  |  | Parallel Channel 6 and Channel 7 Detector Signals
Video Channels 1,3,5, and 7 are fast (500 Kpix/chan/sec) channels
Video Channels 0,2,4, and 6 are slow ( 50 Kpix/chan/sec) channels
Clock & Bias Board |  |  | SN 34
DAC Reference Strapping
Signal |  |  | Value | Jumper | Position
BIASVREF0 |  |  | Bipolar | JP3 | 2-3
BIASVREF1 |  |  | Bipolar | JP1 | 2-3
FBIASVREF |  |  |  | JP2 | 1-2
CLKVREF0 |  |  | Bipolar | JP5 | 2-3
CLKVREF1 |  |  | Bipolar | JP6 | 2-3
CLKVREF2 |  |  | Bipolar | JP4 | 2-3
Power Supply Selection
Signal |  | Value | Jumper | Position
+5VDIN Enable |  | Disable | JP10 | Open
+1.8V Src Select |  | External | JP9 | 1-2
Tst Osc Enable |  | Disable | JP11 | 1-2
AGND Strap 1 |  | In | JP8 | Strapped
AGND Strap 2 |  | In | JP7 | Strapped
Master Control Board |  |  | SN 36
Jumper positions
Signal |  | Value | Jumper | Position
SER CFG WRT EN |  | Enable | JP1 | 2-3
INT/ EXT SYNC SRC |  | Int Sync | JP2 | OPEN
INT / EXT CLK SRC |  | Int Clk | JP3 | 1-2
FREQ RANGE SLCT |  | 25-50Mhz | JP4 | OPEN
3.3VD PWR SLCT |  | External | JP5 | 2-3
IGNORE FC |  | Cont Enbl | JP6 | 2-3
CLK CONFIG[0] |  | Full Clk | JP7 | 2-3
CONVERT SYNC |  | No Conv | JP8 | 2-3
ENABLE CRC |  | Enable | JP9 | 2-3
CLK CONFIG[1] |  | Full Clk | JP10 | 2-3
OSC ENABLE |  | Enable | JP11 | OPEN

## Buy List
Item | Quantity | Manufacturer | Part Number | Supplier | Supplier Part number | Description | Unit Price
1 | 2 | Amphenol | PT06A-20-41S-SR | Mouser | 654-PT06A-20-41S-SR | Cryostat => Cable Connector | 80.25
1 | 2 | Amphenol | PT06A-20-41P-SR | Mouser | 654-PT06A-20-41P-SR | Cryostat => Cable Connector | 63.45
3 | 2 | Amphenol | PT02A-20-41S | Mouser | 654-PT02A-20-41S | DHE Bulkhead Connector | 56.06
4
5
6
7
8
9
10
11
12
13
14
15
16
17
18
19
20
21
22
23
24
25
26
27